## Slide 1
Principles of EQA
Annette Thomas
Chair IFCC C-AQ
Director WEQAS
www.weqas.com

## Slide 2: What is External Quality Assessment?
proficiency testing scheme
System designed to objectively assess the quality of results by an external agency.

## Slide 3: IQC vs EQA
 | Internal Quality Control | External Quality Assurance
Results | Known | Unknown
Results Available | Immediately | Later, when report received
Frequency of Testing | Minimum daily, per batch, per shift | Periodically eg monthly/ Bimonthly quarterly
Concentrations | Normal, abnormal | Pathological range covered over multiple samples
Assess | Imprecision | Accuracy & imprecision
Comparison | Your lab only | Your lab to all labs & other labs using your method

## Slide 4: Objectives of EQA
Provide a measure of the quality of a test
To supplement internal quality control procedures
Provide a measure of the “state of the art” of a test
To obtain consensus values when true values are unknown
To investigate factors in performance (methods, staff etc)
To act as an educational stimulus to improvement in performance
To provide a Post market vigilance service
To provide evidence and monitoring of harmonisation strategies
IFCC 1977

## Slide 5: General Cycle of EQAS
User analyses EQA samples
Results sent to EQA organiser
EQA report received
EQA report reviewed
Corrective action taken if required
EQA samples received

## Slide 6: Planning an EQA?
Design for clinical need
Material type & frequency
  Identify performance specification
  Data analysis
  Reports
  Education element
  Troubleshooting support

## Slide 7: EQA for diabetes programme What tests do we need to assess?
Analyte | Material | Methods
Glucose / Lipids | Serum | Lab & POCT
HbA1c | Whole blood | Lab & POCT
Urine Albumin | Urine | Lab & POCT

## Slide 8: Material & Frequency
Appropriate matrix
Stable, homogeneous material
Appropriate concentration range
Challenging samples
Appropriate frequency of testing

## Slide 9: Sample Matrix
  Good long term stability
  Very easy to transport
  Commutability issues
  Expensive
  Difficult to prepare - reconstitution errors
  Denaturation of proteins/lipoproteins
  Issues with assigning reference targets
Lyophilized material – serum / urine whole blood

## Slide 10: Sample Matrix
  Gold standard
  Commutable – performs as the patient samples
  Easy to use
  Can set reference targets
  Very expensive
  Difficult to obtain different concentration levels
  Unstable analytes
  Storing and transport very expensive
  Risk of bacterial contamination
Fresh/ Frozen clinical material - from individual donors or pool of donors

## Slide 11: Appropriate Concentration Levels
Range of concentration levels
Not just around the reference interval
Cover the analytical and pathological range
Challenging samples

## Slide 12: Planning an EQA?
Design for clinical need
Material type & frequency
Identify performance specification
  Data analysis
  Reports
  Education element
  Troubleshooting support

## Slide 13: Target Value
Determining the ‘right’ , ‘true’, ‘correct’ value?

## Slide 14: Target Value
Reference value
  “True” value with meteorological traceability.
  Commutable EQA material assayed using:
  Reference methods eg HbA1c
  By an Accredited reference laboratory
    Using higher order methods eg IDGC-MS glucose, creatinine.
    Certified Reference material

## Slide 15: Using Reference Method target

## Slide 16: Target Value
Statistical
  Overall mean, median
  Method, instrument mean
  Shows the ‘state of art’ of method performance but gives no indication of the “true” value

## Slide 17: Using Statistical Comparison
Mean ± 2 SD = 95% returned results are acceptable
This gives information on whether your result is acceptable compared with all results and your method
But does it give information on whether your result is acceptable compared to the ‘true’ patient result?
Would your result alter the clinical management of the patient?
Does not provide context to clinical utility

## Slide 18: External Quality Assurance Programs
Use quality standards to allow labs to assess their performance and respond accordingly
the quality standard is the allowable difference from a target
A tool for review of QAP results
Can be based on
  statistical comparison
  expert opinion
  clinical need
  other criteria

## Slide 19: Target values used in Quantitative EQA
Reference values
Gold standard
Gravimetric
Overall mean / median
Used if no ref and data Gaussian.
Method mean / median
Peer group assessment only
Analyser mean
Peer group assessment only.
Loss of information for Assessment of accuracy

## Slide 20: What is the best approach?
Fresh Patient / commutable samples using reference method as target
Fresh patient / commutable samples using a secondary reference method as target
Fresh patient / commutable samples using data from peer group of expert laboratories as target
Mean / median of method group with acceptable performance using fresh patient / commutable samples
Commercial material. - It is essential to understand how the matrix affects the performance of methods if non commutable material is used. Only peer review within instrument / method group can be undertaken in these instances.

## Slide 21: Determining Performance specifications.

## Slide 22: Acceptable Performance Specification
A range of values around the target value that is considered acceptable laboratory performance
A tool for review of EQA results
It provides a simple tool to allow a rapid, standardised assessment of EQA results in both numerical and graphical report formats.
A result outside the acceptable range should alert the laboratory that that their assay may produce results that are at risk of detrimentally affecting clinical decision making.

## Slide 23: Using Statistical Comparison
Mean ± 2 SD = 95% returned results are acceptable
This gives information on whether your result is acceptable compared with all results and your method
May not be clinically relevant
TEe = 10%
Method 1
Method 2
TEa = 5%

## Slide 24: Which Performance Criteria to use?
 | Regulatory requirement, TE |  | Acceptable performance from EQA, TE % |  | 
Test | CLIA (US) | Rilibak (Germany) | RCPA (Aus) | WEQAS (UK) | Labquality (Finland)
Glucose | ± 0.33 mmol/l or ± 10% | ± 15% | ± 0.5 mmol/l or ± 10% | ± 7.5%@4 mmol/l | ± 6.0%
HbA1c | n/a | ± 18% | ± 0.5 or ± 5% | ± 7% | 

## Slide 25: Specification Hierarchy
Analytical performance specification based on clinical outcomes
Analytical performance specification based on biological variation
“State of the art” - Interlaboratory variation
What we need but data not readily available
Data available but not always achievable
What we can achieve but may not be “fit for purpose”
Improvements in methods / technology
Data from outcome studies
Model 1
Model 2
Model 3

## Slide 26: From Biological goals
Test | I (%) | B (%) | TE (%) (0.01)
Glucose | 2.2 | 1.9 | 7.0
HbA1c | 1.7 | 1.5 | 5.5
Desirable quality specification can be calculated from:
I < 0.5CVw
B< 0.25 (CVw2 + CVb2)½
TE = 2.33 I + B (a<0.01) for EQA

## Slide 27: Analytical goals
 |  | Biological goals |  |  | WEQAS TE criteria |  | 
Analyte | Conc. | I (%) | B (%) | TE (0.01) | SD | 2 SD | TE
Albumin | 40 | 1.6 | 1.3 | 4.9 | 1.3 | 2.6 | 6.5
Bicarb | 20 | 2.4 | 1.6 | 7.2 | 1.3 | 2.6 | 13.0
Ca | 2.3 | 1 | 0.8 | 3.1 | 0.05 | 0.1 | 4.3
Cl | 100 | 0.6 | 0.5 | 1.9 | 1.4 | 2.8 | 2.8
Creat | 80 | 2.2 | 3.4 | 8.4 | 8 | 16 | 20.0
Glucose | 4.2 | 2.2 | 1.9 | 7.0 | 0.16 | 0.32 | 7.6
Mg | 0.8 | 1.8 | 1.8 | 6.0 | 0.03 | 0.06 | 7.5
Osmo | 245 | 0.7 | 0.4 | 2.0 | 3.4 | 6.8 | 2.8
Phos | 0.8 | 4.3 | 3.2 | 13.1 | 0.03 | 0.06 | 7.5
K | 4 | 2.4 | 1.8 | 7.4 | 0.08 | 0.16 | 4.0
Na | 135 | 0.4 | 0.3 | 1.2 | 1.5 | 3 | 2.2
T P | 70 | 1.4 | 1.2 | 4.4 | 1.6 | 3.2 | 4.6
Urate | 0.34 | 4.3 | 4.8 | 14.8 | 0.02 | 0.04 | 11.8
Urea | 8 | 6.2 | 5.5 | 19.8 | 0.35 | 0.7 | 8.8
Highlighted TE are those where Biological goals not achievable

## Slide 28
Chart: Creatinine Precision Profile (CV %)
2011 | 25.4087 | 3.8661 | 3.4116 | 4.1388 | 3.0432 | 3.0993 | 3.3484 | 9.2824 | 3.9237 | 5.6385 | 3.4944 | 3.9558 | 3.0965 | 3.4166 | 5.4516 | 9.5035 | 25.0436 | 3.9756 | 3.9977 | 4.0911 | 4.0942 | 3.1386 | 3.1109 | 4.0232 | 3.2709 | 23.0611 | 3.7202 | 4.5772 | 3.8433 | 3.0866 | 3.5494 | 5.6159 | 9.3475 | 3.9056 | 3.8595 | 3.9363 | 3.1133 | 3.6133 | 3.6934 | 21.5348
2012 | 17.6992 | 3.7907 | 4.2462 | 3.5007 | 7.2546 | 3.0384 | 3.1807 | 3.1947 | 4.3965 | 3.5197 | 2.9043 | 2.7283 | 5.187 | 3.4323 | 2.9164 | 3.6694 | 6.7943 | 3.8344 | 3.5975 | 4.1938 | 14.4429 | 3.326 | 3.493 | 2.8297 | 15.8711 | 4.4697 | 3.532 | 2.8943 | 3.5155 | 3.6668 | 29.2311 | 4.711

## Slide 29: Planning an EQA?
Design for clinical need
Material type & frequency
Identify performance specification
Data analysis
Reports
  Education element
  Troubleshooting support

## Slide 30: What to Look for in the EQA Reports
Design should be such that the user can:
  Understand the report
  take appropriate action
  provide help and improvement
  Is education

## Slide 31: What to Look for in EQA Reports
What is the turn around time?
  how long after the due date will you receive your report?
  the earlier the better
  can remember what was happening at that time
  can take corrective action sooner
Is the report electronic or paper?
  Electronic quicker but not all sites may have printer access.

## Slide 32: What to Look for in EQA Reports?
Report should
  Clearly identify performance
    Assess accuracy to target value
    Assess imprecision
    Assess linearity
    Assess performance over time
    Identify errors and poor performance
  Clearly identify method / analyser performance

## Slide 34
Example of performance alert.

## Slide 35: Planning an EQA?
Design for clinical need
Material type & frequency
Identify performance specification
Data analysis
Reports
Education element
  Troubleshooting support

## Slide 36: Educational role
Pre-analytical effects
Number and type of methods used
Performance of methods used
    accuracy
    precision
Susceptibility of methods to interference
    including other analytes and matrix
Interpretation of results

## Slide 37: Educational role
Education on methods and interferences
eg
  Concentration near diagnostic “cut off”
  Samples with known variants for HbA1c
  Samples with added interferances

## Slide 38: Intralaboratory variation – Roche TnT
Chart: Troponin T (ng/L)
| 3.5245000000000006 | 13.9215 | 22.6845 | 38.724500000000006
AF | 6.0402 | 2.8702 | 1.6487 | 1.405
AJ | 5.4128 | 1.4762 | 1.4607 | 0.90206
AQ 1 | 12.2107 | 4.8275 | 3.0459 | 1.4768
AQ 2 | 8.6062 | 3.5115 | 3.2586 | 2.4751
AH 1 | 4.4135 | 2.8361 | 2.034 | 1.9454
AH 2 | 5.4429 | 3.1323 | 2.0459 | 2.02
AI 1 | 6.9075 | 2.7465 | 2.0206 | 1.6578
AI 2 | 2.034 | 1.1036 | 1.3484
JU 1 | 5.6531 | 1.8945 | 2.1056 | 1.4207
JU 2 | 4.2733 | 1.6774 | 1.4992 | 0.89761

## Slide 39: Interferences - Effect of icterus on assays
A pre-analytical EQA Scheme for sample integrity:
A Weqas study to monitor the effectiveness of Serum Indices

## Slide 40: Planning an EQA?
Design for clinical need
Material type & frequency
Identify performance specification
Data analysis
Reports
Education element
Troubleshooting support

## Slide 41: Troubleshooting Support
Can you provide a consultation service?
Available by phone, fax, e-mail
Provide help with
  method classification queries
  analyte problems
  report interpretation
  troubleshooting

## Slide 42
Problem Solving Flow Chart
Pages 18-19 of SP-QL1-IntLabEQA

## Slide 43
[3] Check whether the
cause is curvilinear data
(m,c,Sy.x,r)
[4] Then the error is random,
check whether there is
clerical error.
[5]
Check for causes
of imprecision
e.g. inexperienced
operators, faulty
equipment, inappropriate
methods
Eliminate
blunder
go to [2]
go to [2]
YES
NO
NO
NO
[1] Are you satisfied with
your imprecision values?
(Sy.x, r)
Problem Solving Flow Chart
IMPRECISION
Pages 18-19 of SP-QL1-IntLabEQA

## Slide 44: Problem Solving Flow Chart
Pages 18-19 of SP-QL1-IntLabEQA

## Slide 45
Curvilinear
(m, c, Sy.x, r)
proportional
(m)
mixed
(m, c)
constant
(c)
Check for time
expired reagents
Is it a “one
point” cal?
Replace
reagents
Check all
calibrators
inc. zero
Check zero
(reagent blank,
serum blank,
instrument zero).
Recalibrate
Run linearity check
Are you satisfied
with slope?
Are you satisfied
with intercept?
Problem Solving Flow Chart
Adjust slope
value of instrument
Recalibrate,
run linearity check.
Check method
specificity
YES
YES
YES
YES
NO
NO
NO
NO
INACCURACY
Pages 18-19 of SP-QL1-IntLabEQA

## Slide 46: EQA and User in partnership
Part of Quality Improvement
Should not be viewed as a pass/fail exercise
Educational – troubleshooting, recommendations of best practice
Identify poor methods
Provide training and help